## Sheet1
Sl. No. | Name Of Company | Target Segment - I/II/III/IV | Approved Eligibile Product Name | Version/ Model name | Specification of each Version/ Model name | Website of Company | Broucher and link of product details
1 | Wipro GE Medical Device Manufacturing Private Limited | II | Revolution Aspire | Revolution Aspire | Generator - 42 KW , Rotation Time - 0.8 Sec, Current Out Put - upto 350 mA, 3.5 MHU X Ray tube with anode heat dissipation rate of 820 KHU/min, Clarity Panel segmented panel 10 mm detector capable of covering 21 mm per second supported by pitch factor of 1.675, Pitch booster technology, Ultrakernel for small anatomy, minimum slice thicknes of 0.625mm , UP/Down Patient Table with load bearing capacity of 180 KGs,Advanced applications like metal artifact reduction (Smart mAR) technique, auto bone Express , vessel IQ express, Virtual colonoscopy , CT Perfusion, Denta Scan etc Advanced sequential iterative reconstruction ( ASIR), Image recon speed of upto 22 fps | https://www.gehealthcare.in | Brochure Attached
2 | Wipro GE Medical Device Manufacturing Private Limited | II | Revolution Aspire | Revolution Aspire Select | Generator - 42 KW , Rotation Time - 0.8 Sec, Current Out Put - upto 350 mA, 3.5 MHU X Ray tube with anode heat dissipation rate of 820 KHU/min, Clarity Panel segmented panel 10 mm detector capable of covering 21 mm per second supported by pitch factor of 1.675, Pitch booster technology, Ultrakernel for small anatomy, minimum slice thicknes of 0.625mm , FIxed Config Patient Table with load bearing capacity of 180 KGs,Advanced applications like , auto bone Express , vessel IQ express, Virtual colonoscopy , CT Perfusion, Denta Scan etc Advanced sequential iterative reconstruction ( ASIR), Image recon speed of upto 22 fps | https://www.gehealthcare.in | Brochure Attached
3 | Wipro GE Medical Device Manufacturing Private Limited | II | Revolution Maxima | Revolution Maxima | Generator - 72 KW , Rotation Time - 0.35 Sec, Current Out Put - upto 600 mA, 7 MHU X Ray tube with anode heat dissipation rate of 1070 KHU/min, 40 mm detector, Patient table weight bearing capacity of 227 KG,minimum slice thicknes of 0.625mm , 0.28Mmm spatial Resolution, Fast Recon speed of 55 fp/s, Advanced applications like , AI based auto positioning, snapshot freeze, 5 beat cardiac CT, ASiR-V for up to 82% dose reduction, FastStroke, Autobone, Vessel IQ,Ultra Kernel, Smart MAR, CT Perfusion 4D, Hepatic VCAR, Thoracic VCAR, Stroke VCAR, TAVI Analysis with Multi-Modality Workstation | https://www.gehealthcare.in | Brochure Attached